Bpharma

Advanced Pharmaceutical Industries Co. Ltd · Jeddah, Kingdom of Saudi Arabia

GxP Risk Assessment

# GxP Risk Assessment

| Document Number VAL/LIMS/GRA/001 | Document Type GRA | Version 0.1 DRAFT | Effective Date Pending approval |
| --- | --- | --- | --- |
| System B-Pharma LIMS (desktop edition) | GAMP Category Category 5 — Custom Application |

Note DRAFT. No approval signatures have been applied. Requirement identifiers in this package were created by this package (no approved URS existed); the System Owner must review and approve them. Technical execution recorded here was performed by an AI assistant (Claude); a qualified human reviewer must review the evidence and countersign.

# Contents

Contents

1. Method

2. GAMP category

3. Assessment

4. Consequence for testing

Approval

# 1. Method

RPN = Severity x Likelihood x Detectability (each 1-5; Detectability 5 = hard to detect). RPN >= 50 High (OQ plus PQ), 15-49 Medium (OQ), < 15 Low. Scores are the author's proposal and require QA confirmation.

# 2. GAMP category

Category 5.

# 3. Assessment

| Area | S | L | D | RPN | Level | Rationale |
| --- | --- | --- | --- | --- | --- | --- |
| Authentication & session | 5 | 3 | 2 | 30 | Medium | A forged or unauthenticated session allows any record to be changed under another identity. |
| Authorisation & workflow | 5 | 3 | 3 | 45 | Medium | Wrong role or skipped state could release an unapproved batch; a later independent review might catch it. |
| Electronic signatures | 5 | 3 | 2 | 30 | Medium | Signature without genuine re-authentication or with unverifiable integrity is legally void under Part 11. |
| Audit trail integrity | 5 | 3 | 1 | 15 | Medium | Silent loss or rewrite of history; structural controls (hash chain, triggers) make it detectable, anchors cover owner-level rewrite. |
| Desktop deployment & secrets | 4 | 3 | 3 | 36 | Medium | Credentials reachable from the shared Windows session; local administrator can bypass application controls. |
| Backup & recovery | 5 | 3 | 5 | 75 | High | Loss of the single PC database is undetectable until needed; no function exists. |
| Non-core modules (ELN, OOS, stability, environmental, reagents, instruments, training, reports) | 4 | 3 | 3 | 36 | Medium | Not individually qualified in this release (see DEV-04). |
| Results, raw data and calculations | 5 | 3 | 2 | 30 | Medium | A wrong or unattributable value reaches a certificate; versions, hashes, review gates and server calculation make change visible. |
| Instrument capture and machine integration | 4 | 3 | 2 | 24 | Medium | Machine-entered values bypass the bench; acceptance by a named analyst, scopes and idempotency limit that. |
| Certificates, OOS and quality records | 5 | 3 | 2 | 30 | Medium | A certificate or a closure that misstates compliance releases a batch wrongly. |
| Controlled documents, training and custody | 4 | 3 | 3 | 36 | Medium | Uncontrolled documents or unqualified analysts; unreceived or overdrawn samples. |
| Reagents, stability and environmental monitoring | 4 | 3 | 3 | 36 | Medium | Expired material or missed time points go unnoticed; the programmes compute them without user action. |
| Audit trail hardening and validated scope | 5 | 3 | 2 | 30 | Medium | Rewriting history as the database owner; a module used without qualification. |
| SPC and trending | 3 | 3 | 2 | 18 | Medium | A wrong control limit or a missed rule violation hides a drift; the output is advisory and writes nothing to a result, sample or release. |
| Directory sign-on and signatures | 5 | 3 | 2 | 30 | Medium | A mis-mapped group or a weak directory connection grants a role or a signature to the wrong person; verified TLS, QA-signed mappings and the local lock limit it. Not tested against a real Active Directory. |
| Periodic review, retention and archive | 5 | 3 | 2 | 30 | Medium | An archive that does not verify, or records that change after archiving, loses the legal record. |
| Global search | 3 | 3 | 2 | 18 | Medium | A role sees a record it may not open; groups are filtered by the same permissions as the screens. |
| Arabic/English interface and formatting | 4 | 3 | 3 | 36 | Medium | A misread decimal or an ambiguous date changes a result or a record; one parsing rule on the screen and the server, fixed formats, Arabic wording is a draft. |
| Installer and release integrity | 5 | 3 | 3 | 45 | Medium | An installer that does not match its manifest cannot be trusted, and a broken one stops all work; the check runs at every start and on a Windows runner before publishing. |
| Built-in database | 4 | 3 | 3 | 36 | Medium | A database without least privilege or supervision loses data or exposes it to other accounts; not exercised on Windows and not yet bundled. |

# 4. Consequence for testing

- Authentication & session: Medium - OQ

- Authorisation & workflow: Medium - OQ

- Electronic signatures: Medium - OQ

- Audit trail integrity: Medium - OQ

- Desktop deployment & secrets: Medium - OQ

- Backup & recovery: High - OQ + PQ

- Non-core modules (ELN, OOS, stability, environmental, reagents, instruments, training, reports): Medium - OQ

- Results, raw data and calculations: Medium - OQ

- Instrument capture and machine integration: Medium - OQ

- Certificates, OOS and quality records: Medium - OQ

- Controlled documents, training and custody: Medium - OQ

- Reagents, stability and environmental monitoring: Medium - OQ

- Audit trail hardening and validated scope: Medium - OQ

- SPC and trending: Medium - OQ

- Directory sign-on and signatures: Medium - OQ

- Periodic review, retention and archive: Medium - OQ

- Global search: Medium - OQ

- Arabic/English interface and formatting: Medium - OQ

- Installer and release integrity: Medium - OQ

- Built-in database: Medium - OQ

# Approval

| Role | Name / Title | Signature | Date |
| --- | --- | --- | --- |
| Prepared by (Validation) |  |  |  |
| Reviewed by (System Owner) |  |  |  |
| Reviewed by (QA) |  |  |  |
| Approved by (QA Head) |  |  |  |