Bpharma

Advanced Pharmaceutical Industries Co. Ltd · Jeddah, Kingdom of Saudi Arabia

Validation Plan

# Validation Plan - B-Pharma LIMS Desktop

| Document Number VAL/LIMS/VP/001 | Document Type VP | Version 0.1 DRAFT | Effective Date Pending approval |
| --- | --- | --- | --- |
| System B-Pharma LIMS (desktop edition) | GAMP Category Category 5 — Custom Application |

Note DRAFT. No approval signatures have been applied. Requirement identifiers in this package were created by this package (no approved URS existed); the System Owner must review and approve them. Technical execution recorded here was performed by an AI assistant (Claude); a qualified human reviewer must review the evidence and countersign.

# Contents

Contents

1. Purpose

2. Scope

3. References

4. System description

5. GAMP category

6. Validation approach

7. Roles

8. Deliverables

9. Acceptance criteria

10. Deviation management

11. Change control

12. Training

13. Periodic review

14. Approval

# 1. Purpose

This plan defines how the B-Pharma LIMS desktop application is validated for GxP use following GAMP 5 (2nd Ed.), 21 CFR Part 11 and EU Annex 11, and states what was and was not verified in this release.

# 2. Scope

- In scope: authentication, sample workflow, electronic signatures, audit trail, desktop packaging and secret handling (URS-001 to URS-050), the release R2 modules (URS-060 to URS-119): master data, methods and calculations, result entry, raw data, instrument capture, the machine API, events, certificates, OOS, quality records, documents and training, custody, reagents, stability, environmental monitoring, worklists, turnaround, audit hardening and the validated-scope register.

- Release R3 (URS-120 to URS-145): SPC and trending, directory (LDAP/AD) sign-on, periodic review, retention and archive, global search, the Arabic and English interface, installer and release integrity, and the built-in database.

- Without OQ in this release (see DEV-04): ELN, the AI assistant, voice input and the camera barcode scanner. QA declares the validated scope of each catalogued module in the application (URS-116).

- Out of scope: network and PostgreSQL server qualification (site IT programme). The Windows installer is exercised only by the CI smoke test (DEV-2026-10-26-01); the target PC is qualified by IQ-W.

# 3. References

- GAMP 5, 2nd Edition

- 21 CFR Part 11; EU GMP Annex 11

- Project directives: AGENTS.md

- docs/gxp-audit-findings.csv (code review findings GX-01 to GX-17)

# 4. System description

A React/TypeScript interface served by an embedded Express server on the loopback interface, packaged with Electron, using a PostgreSQL database on the same PC. Audit entries form a SHA-256 hash chain; audit and signature tables are append-only through database triggers.

# 5. GAMP category

Category 5 - custom application. Infrastructure components (PostgreSQL, Electron, Node.js) are Category 1.

# 6. Validation approach

V-model: URS to FRS, GRA sets rigor, IQ/OQ/PQ, RTM ties them together. Prior review and fixing work (GX series) is supporting history; formal evidence is the execution recorded in the OQ/PQ/IQ documents, generated from the evidence files in validation/evidence/ (tests/security-controls.sh, tests/backup-restore.sh, tests/desktop-smoke.cjs, the Vitest JUnit file and tests/e2e/run.cjs) executed on 2026-10-26. The RTM is generated from those files and node validation/build/check.cjs fails when a Critical or High requirement has no passed test and no open deviation.

# 7. Roles

| Role | Responsibility | Name |
| --- | --- | --- |
| System Owner | Approves URS, accepts residual risk |  |
| QA | Approves package, reviews deviations |  |
| Validation Lead / IT | Executes IQ on the target PC, owns backups |  |
| Technical execution | Automated OQ/PQ run (AI assistant, Claude) | Claude (AI) |
| End users | Perform and observe PQ |  |

# 8. Deliverables

| Document | Number | Purpose |
| --- | --- | --- |
| Validation Plan | VAL/LIMS/VP/001 | This document |
| URS | VAL/LIMS/URS/001 | Requirements |
| GxP Risk Assessment | VAL/LIMS/GRA/001 | Risk-based rigor |
| FRS | VAL/LIMS/FRS/001 | How requirements are met |
| RTM | VAL/LIMS/RTM/001 | Traceability (also RTM.csv) |
| IQ | VAL/LIMS/IQ/001 | Installation |
| OQ | VAL/LIMS/OQ/001 | Function |
| PQ | VAL/LIMS/PQ/001 | Business scenarios |
| DIA | VAL/LIMS/DIA/001 | ALCOA+ |
| VSR | VAL/LIMS/VSR/001 | Summary and conclusion |

# 9. Acceptance criteria

- Every Critical URS has an executed Pass (P) or a QA-accepted P* justification.

- All deviations CLOSED or risk-accepted in writing before release to GxP use.

- IQ executed on the target PC; end-user PQ completed.

# 10. Deviation management

Deviations are numbered DEV-YYYY-MM-DD-NN, recorded in the VSR and deviations.csv with root cause, correction and status.

# 11. Change control

Any change after approval follows the site change-control SOP; tests/security-controls.sh and the RTM are re-run, and the build is identified by its git commit and file hashes (IQ).

# 12. Training

Users complete the in-application software training before applying signatures (URS-024). PQ should be performed by representative end users.

# 13. Periodic review

At least annually, after any change to audit/signature code, and after any restore from backup.

# 14. Approval

| Role | Name / Title | Signature | Date |
| --- | --- | --- | --- |
| Prepared by (Validation) |  |  |  |
| Reviewed by (System Owner) |  |  |  |
| Reviewed by (QA) |  |  |  |
| Approved by (QA Head) |  |  |  |